# Incident Synopsis

A researcher with approximately three years of laboratory experience working with hydrogen storage materials was working in the laboratory at approximately 7:30 p.m. on April 19, 2005.  Several other researchers were working in adjacent laboratories.

The researcher had prepared a sample of aluminum deuteride, AlD3, (D represents a heavier isotope of hydrogen with an extra neutron in the nucleus, but which is chemically the same) by reacting lithium aluminum deuteride and aluminum chloride in diethyl ether.  The actual composition/phase of the material synthesized was unknown, but the researcher had attempted to produce the gamma phase of aluminum deuteride.  The synthesis steps used to produce the material were complete and the researcher attempted to seal the material in a glass ampoule for offsite shipment and analysis.  The sample size was approximately 1 gram.

The ampoule with the sample had previously been placed under vacuum and had been isolated from the atmosphere.  The process for sealing the ampoule was to first place the sample-containing ampoule in a liquid nitrogen bath to cool the sample down to near 77K.  The ampoule was then removed from the liquid nitrogen and a torch was used to melt the neck of the ampoule to seal it.  The ampoule was slowly rotated while heat was applied to the neck of the ampoule.  Typically, this results in collapse of the ampoule, sealing the sample.  However, in this instance, a bubble formed where the heat was applied and a hole formed in the ampoule.  This allowed air to enter the previously evacuated container.

After about 30 seconds the ampoule “exploded” and glass from the container was sprayed outward.  Some of the shards embedded into the researcher’s arms, face and torso (the researcher was wearing safety glasses, but no other safety equipment).  Other researchers working in an adjacent laboratory heard the noise and came to see what had happened.  The researcher was taken to the hospital and treated.  Larger shards of glass were removed and the injured area was cleaned in an effort to remove the smaller pieces of glass. The researcher stayed home the next day, but returned the second day after the incident.  There appears to be no permanent damage to the researcher and he/she is continuing work in the laboratory.  Some of the glass pieces have remained in the researcher’s arm and are expected to slowly work to the skin surface where they will fall out or be removed.

# Cause

Clearly, penetration of the glass ampoule and subsequent ingress of air and water vapor resulted in a rapid chemical reaction resulting in the energetic destruction of the sample container.  Furthermore, the 30 second delay suggests that temperature changes played an important role.  There are three potential sources for a high overpressure event to occur in this scenario, (1) detonation of the sample material by a rapid reaction, (2) rapid over pressurization of the glass ampoule (without ignition) by hydrogen released from the sample and, (3) release of hydrogen from the sample and subsequent detonation of the hydrogen/air mixture in the sample tube.  The specific source or reaction chain that produced the event, however, can not be identified because the composition/phase(s) of the sample material at the time of the event are not known and, furthermore, there was no mention by the researcher of a bright flash of light at the time of the event.

Previous experience with the expected synthesized material, , , or  phase AlD3, suggests that it is stable when exposed to air at room temperature.  Furthermore, aluminum hydride (AlH3) is known to be metastable at room temperature; that is, it will not decompose even though such decomposition is thermodynamically favorable.  Recent work suggests that AlD3 is also metastable at room temperature.  Since it is unlikely that the sample temperature reached much higher than ambient temperature during the 30 second interval, thermal decomposition of the phase would not be expected to occur.  On the other hand, the phases are all thermodynamically unstable, so any reactant, such as water, could release the hydrogen by forming an oxide phase of Al under the proper conditions.

There may also be other materials present that could act as precursors to a reaction resulting in the release of hydrogen, or they may even be the main source of energy release.  For example, incomplete or failed synthesis could have resulted in lithium aluminum deuteride (LiAlD) in the finished product.  The hydride form of the material is well known to be pyrophoric and water reactive.  Any remaining aluminum chloride might also play a role.

There also appears to have been sufficient hydrogen (~0.1 mol) contained in the ~1 gm of sample material that, if rapidly released, could have been the cause.   It is estimated that the hydrogen gas pressure within the ampoule could have reached ~200 psi if it were released sufficiently fast to overcome the loss through the hole caused by the torch.  At this pressure, the glass could have failed even without ignition.  Certainly, detonation of a hydrogen/air mixture in the glass ampoule could have occurred if hydrogen were released sufficiently rapidly to reach the concentration level required for detonation or deflagration.

# Contributing Factors

# Like many accidents, a number of factors contributed to this incident:

1.  The material at various stages of the synthesis process was, in all likelihood, not the material it was presumed to be.  As described above, large amounts of precursor material may have been present; the deuterated alane may have been of a different phase and/or of unexpectedly high reactivity. It may not have been completely desolvated.

2.  The material that was to be synthesized was presumed to have a low      reactivity based upon historical and literature information. This is not a valid    presumption.

- Laboratory procedures did not account for safeguards against unforeseen hazards such as the possibility of the presence of a highly reactive material.

Lessons Learned

Lessons learned beyond common laboratory practices, e.g. importance of wearing safety glasses, utilizing clothing such as lab coats, not working alone when performing potentially hazardous operations, etc. include:

- Metal hydride materials of a composition which is not well characterized should be handled with procedures that assume a “worst case” for that class of materials, intermediates or precursors.

- Laboratory procedures should be in written form and should be adopted only after performing a safety vulnerability analysis and adopting appropriate risk mitigation steps.

- Working with small amounts of material does not provide assurance of safety.

- The method described to seal samples that are highly reactive upon exposure to air is not recommended.  An alternative method used by the laboratory for packaging aluminum hydride samples for offsite shipment is as follows:  Aluminum hydride powders (0.5 g - 1.0 g) are sealed in glass bottles with a cap. The bottles are sealed in a thick plastic bag under Ar using a plastic bag sealer in the glove box. The sample bags are then removed from the glove box and sealed under vacuum in a much larger plastic bag using a vacuum sealer.  The purpose of the large evacuated bag is to contain any evolved hydrogen gas if the material begins to decompose and the internal bottle & bag rupture.  It is a good idea to calculate the maximum volume of H2 (at 1-atm) that could be released by the sample to determine if the outer bag is sufficiently large to contain the evolved gas.  The package is then inserted into a cardboard box of sufficient volume (about 1-liter) to accommodate the expanded outer bag.